## Series names
Series description match | Series description exclude | Vendors | Standardized name (blank = ignore) | Notes
3D_TOF |  | Philips, Siemens
TRANCE |  | Philips
LOCALISER
SURVEY |  | Philips
DEMO |  | Philips
PCASL | SOURCE | Philips |  | Ignore any PCASL scan without SOURCE in description
calibration |  | GE
Cor_B0map |  | Philips, GE | B0
T2star_HIGHRES |  | Philips | T2STAR
Cor_T1map_resptrig | FLIPPED | Philips | T1MAP
Cor_T1map_resptrig_flipped |  | Philips | T1MAP_FLIPPED
Cor_T1MOLLI |  | Philips | MOLLI
Cor_T1_MOLLI |  | Philips | MOLLI
Cor_B1_map_BH |  | GE | B1MAP
Cor_B1map_BH |  | Philips | B1MAP
Cor_DWI_RespTrig_flipped |  | Philips | DWI_FLIPPED
Cor_DWI_RespTrig | FLIPPED | Philips, GE | DWI
PC_BH_Left_Renal_Artery_UKRIN |  | Philips | PC_LEFT
PC_BH_Right_Renal_Artery_UKRIN |  | Philips | PC_RIGHT
Cor_T1W |  | Philips | T1W
Cor_T2W |  | Philips | T2W
Cor_T2_MAPPING |  | Philips | T2_MAPPING
mDIXON |  | Philips | MDIXON
2D_MTR |  | Philips | MTR
Cor_BASE_SE_EPI_RespTrig |  | Philips | BASE
SOURCE-COR_FAIR_SE_EPI_1500 |  | Philips | FAIR_1500
SOURCE-COR_FAIR_SE_EPI_400 |  | Philips | FAIR_400
SOURCE-COR_FAIR_SE_EPI_750 |  | Philips | FAIR_750
SOURCE-COR_FAIR_SE_EPI_1000 |  | Philips | FAIR_1000
SOURCE-PCASL_SE_EPI_1500 |  | Philips | PCASL_1500
Cor_B0map_pCASL_BH |  | Philips | B0MAP_PCASL
AO_flow_NSA1 |  | Philips | AO
SOURCE-PCASL_SE_EPI_500 |  | Philips | PCASL_500
SOURCE-PCASL_SE_EPI_750 |  | Philips | PCASL_750
SOURCE-PCASL_SE_EPI_1000 |  | Philips | PCASL_1000
B0 map |  | Siemens | B0
T2*_high_res |  | Siemens | T2STAR
ep2d_se_T1map_kidney | Rev | Siemens | T1MAP
ep2d_se_T1map_kidney_Rev |  | Siemens | T1MAP_FLIPPED
T1 MOLLI high res |  | Siemens | MOLLI
B1 mapping |  | Siemens | B1MAP
DWI flipped |  | Siemens | DWI_FLIPPED
DWI | FLIPPED | Siemens | DWI
L renal artery |  | Siemens | PC_LEFT
R renal artery |  | Siemens | PC_RIGHT
T1W |  | Siemens | T1W
T2W |  | Siemens | T2W
T2_mapping |  | Siemens | T2_MAPPING
DIXON |  | Siemens | MDIXON
MToff |  | Siemens | MTR_OFF
MTon |  | Siemens | MTR_ON
FAIR_1500
FAIR_400
FAIR_750
FAIR_1000
pCASL 500ms |  | Siemens | PCASL_500
pCASL 750ms |  | Siemens | PCASL_750
pCASL 1000ms |  | Siemens | PCASL_1000
pCASL 1500ms |  | Siemens | PCASL_1500
B0pCASL |  | Siemens | B0MAP_PCASL
Aorta PC-MRI |  | Siemens | AO
Ax_3D_TOF_BH |  | GE
Cor_T2star_BH_highres |  | GE | T2STAR
Cor_T1map_MPMS |  | GE | T1MAP
Cor_T1map_flipped |  | GE | T1MAP_FLIPPED
Cor_T1_MOLLI_BH_HighRes |  | GE | MOLLI
Cor_DWI_RespTrig_b0 |  | GE | DWI_FLIPPED
PC_L_renal_artery_BH |  | GE | PC_LEFT
PC_R_renal_artery_BH |  | GE | PC_RIGHT
Cor_T1weighted_BH |  | GE | T1W
Cor_T2weighted_BH |  | GE | T2W
Cor_T2map |  | GE | T2_MAPPING
Dixon_BH |  | GE | MDIXON
Cor MT off BH MPS |  | GE | MTR_OFF
Cor MT on BH |  | GE | MTR_ON
Cor M0 |  | GE | BASE
Cor FAIR 1500 |  | GE | FAIR_1500
Cor FAIR 400 |  | GE | FAIR_400
Cor FAIR 750 |  | GE | FAIR_750
Cor FAIR 1000 |  | GE | FAIR_1000
Cor pCASL 1500 |  | GE | PCASL_1500
Cor_B0map_BH_no_reshim |  | GE | B0MAP_PCASL
PC_aorta |  | GE | AO
Cor pCASL 500 |  | GE | PCASL_500
Cor pCASL 750 |  | GE | PCASL_750
Cor pCASL 1000 |  | GE | PCASL_1000

## Checks
Series description matchers | Series description exclusion | Vendors | Software Version | DICOM Tag | Parameter | Operator | Expected Value | Read Value | WARNING IF FAIL | Notes
T1map,T1MAP_FLIPPED,BASE,FAIR,PCASL,DWI, DWI_FLIPPED |  | Philips |  | (0018, 1310) | Acquisition Matrix | == | [0, 128, 126, 0]
T1MAP,T1MAP_FLIPPED,T2_MAPPING,BASE,FAIR,PCASL,B0 | B0MAP_PCASL | GE,Siemens |  | (0018, 1310) | Acquisition Matrix | == | [0, 128, 128, 0]
T1MAP,T1MAP_FLIPPED,T2STAR,DWI,BASE,FAIR,PCASL | B0MAP_PCASL | GE,Philips,Siemens |  | (0018, 0050) | Slice thickness | == | 5 |  |  | Note all PCASL are not 'B0map PCASL'
T1MAP,T1MAP_FLIPPED,T2_MAPPING,T2STAR,DWI,BASE,FAIR,PCASL,MDIXON |  | GE,Philips,Siemens |  | (0018, 0088) | Spacing between slices | == | 5.5
T1MAP,T1MAP_FLIPPED,T2_MAPPING,T2STAR,BASE,FAIR,PCASL |  | Philips |  | (2001,1018) | Number of slices | == | 5 |  |  | Check Siemens
T1MAP,T1MAP_FLIPPED,T2_MAPPING,T2STAR,BASE,FAIR,PCASL,B1MAP |  | GE |  | (0021, 104F) | Number of slices | == | 5
T1MAP,T1MAP_FLIPPED,T2_MAPPING,T2STAR,BASE,FAIR,PCASL,MTR,MDIXON | B0MAP_PCASL |  |  | (0018, 1312) | Phase encoding direction | == | ROW
T1MAP,T2_MAPPING,DWI,ASL,BASE,FAIR,PCASL |  |  |  | (0028, 0030) | Pixel spacing | == | [3, 3]
MOLLI,T2STAR,MTR,T1W,T2W |  |  |  | (0028, 0030) | Pixel spacing | == | [1.5, 1.5]
MOLLI, T2STAR,MTR,T2W, T1W |  |  |  | (0018, 1310) | Acquisition Matrix | == | [0, 256, 256, 0] |  |  | It was repeated for almost all sequences from first row...
MOLLI |  | Philips,Siemens |  | (0018, 1310) | Acquisition Matrix | == | [0, 256, 256, 0]
MOLLI |  | GE |  | (0018, 1310) | Acquisition Matrix | == | [0, 224, 256, 0] |  | WARNING | It was repeated for almost all sequences from first row...
T1MAP,T1MAP_FLIPPED,T2STAR,BASE,FAIR,PCASL,MTR,NDIXON |  | Philips |  | (2005, 140f)/(0018, 9069) | Parallel acquisition | == | 2 |  |  | NESTED in 2005,140F
T1MAP,T1MAP_FLIPPED,T2STAR,BASE,FAIR,PCASL,MTR |  | Siemens |  | (0051,1011) | Parallel acquisition | == | p2
T1MAP,T1MAP_FLIPPED,DWI,DWI_FLIPPED,BASE,FAIR,PCASL |  | Philips |  | (2005, 140f)/(0018, 9081) | Partial Fourier | == | YES |  |  | NESTED in 2005,140F
MDIXON,T2STAR,PC_RIGHT,PC_RIGHT |  | Philips |  | (2005, 140f)/(0018, 9081) | Partial Fourier | == | NO
T1MAP,T1MAP_FLIPPED,T2_MAPPING,T2STAR,DWI,BASE,FAIR,PCASL,MTR,MDIXON |  | GE |  | (0018, 0083) | Number of averages (1/partial fourier factor) | == | 1
T1MAP,T1MAP_FLIPPED,T2_MAPPING,BASE,FAIR,PCASL,DWI,DWI_FLIPPED,B0 |  | GE |  | (0028,0010) | Rows | == | 128
MOLLI, T2STAR,MTR,T2W,MDIXON |  | GE |  | (0028,0010) | Rows | == | 256
T1MAP,T1MAP_FLIPPED,T2_MAPPING,BASE,FAIR,PCASL,DWI,DWI_FLIPPED,B0 |  | GE |  | (0028,0011) | Columns | == | 128
MOLLI, T2STAR,MTR,T2W,MDIXON |  | GE |  | (0028,0011) | Columns | == | 256
ASL | PCASL |  |  | (0020, 0105) | Time points | == | 30
T2STAR |  | Philips |  | (0018, 0091) | Number of echoes | == | 12
T2STAR |  | GE |  | (0019, 107E) | Number of echoes | == | 12
T2STAR |  | GE,Philips |  | (0018, 0081) | Echo time | contain | [4.6, 9.2]
T2STAR |  | Siemens |  | (0018, 0081) | Echo time | contain | [4.92, 9.84]
T2STAR |  |  |  | (0018, 1314) | Flip angle | == | 18
T2STAR |  |  |  | (0018, 0080) | TR | == | 80
T2STAR |  |  |  | (0018, 0095) | Bandwidth | range | [200, 400] | 299.6
T2_MAPPING |  | Philips |  | (0018, 0050) | Slice thickness | == | 4.5
T2_MAPPING |  | GE |  | (0019, 107E) | Number of echoes | == | 10
T2_MAPPING |  | Philips |  | (0018, 0091) | Number of echoes | == | 10
T2_MAPPING |  |  |  | (0018, 0081) | Echo time | contain | [12.9, 25.8]
T2_MAPPING |  | GE,Philips |  | (0018, 1314) | Flip angle | == | 90
T2_MAPPING |  | Philips,Siemens |  | (0018, 0080) | TR | range | [3000, 6000]
T2_MAPPING |  |  |  | (0018, 0095) | Bandwidth | range | [200, 400] | 299.6
T2_MAPPING |  | Philips |  | (2001, 105F)/(2005, 107b) | Phase encoding direction | == | RL
T2_MAPPING |  | Philips |  | (2005, 140f)/(0018, 9069) | Parallel acquisition | == | 3
T2_MAPPING |  | Siemens |  | (0051,1011) | Parallel acquisition | == | p3
T2_MAPPING |  | GE |  | (0043, 1083) | Parallel acquisition factors | == | [0.344, 1]
T2_MAPPING |  | Philips |  | (2001, 1018) | Number of slices | == | 5
T1W |  |  |  | (0018, 1314) | Flip angle | == | 42
T1W |  |  |  | (0018, 0080) | TR | == | 155
T1W |  |  |  | (0018, 0095) | Bandwidth | range | [600, 900] | 787.5
T1W |  |  |  | (2005, 140f)/(0018, 9081) | Partial Fourier | == | YES
T1W |  | Philips,GE |  | (0018, 0081) | Echo time | == | 2.3
T1W |  | Siemens |  | (0018, 0081) | Echo time | == | 2.46
T1W |  | Philips |  | (2001,1018) | Number of slices | == | 17
T1W |  | GE |  | (0021, 104F) | Number of slices | == | 17
T1W |  | Philips |  | (2005, 140f)/(0018, 9029) | Oversampling | == | 2D |  |  | NESTED
T1W |  | GE |  | (0043, 1083) | Parallel acquisition factors | == | [0.672, 1]
T2W |  | Philips,GE |  | (0018, 1314) | Flip angle | == | 90
T2W |  |  |  | (0018, 0080) | TR | range | [800, 1000] | 877
T2W |  |  |  | (0018, 0095) | Bandwidth | range | [600, 900] | 787.5
T2W |  |  |  | (2005, 140f)/(0018, 9081) | Partial Fourier | == | YES | 61
T2W |  | Philips,Siemens |  | (0018, 0081) | Echo time | == | 60
T2W |  | GE |  | (0018, 0081) | Echo time | range | [55, 65] | 62
T2W |  | Philips |  | (2001,1018) | Number of slices | == | 17
T2W |  | GE |  | (0021, 104F) | Number of slices | == | 17
T2W |  | Philips |  | (2005, 140f)/(0018, 9029) | Oversampling | == | 2D |  |  | NESTED
T2W |  | GE |  | (0043, 1083) | Parallel acquisition factors | == | [0.344, 1]
DWI,DWI_FLIPPED |  |  |  | (0018, 1314) | Flip angle | == | 90
DWI,DWI_FLIPPED |  | Philips,Siemens |  | (0018, 0080) | TR | range | [3000, 6000] | 3000
DWI,DWI_FLIPPED |  |  |  | (0018, 0095) | Bandwidth | range | [1800, 2200] | 1988
DWI,DWI_FLIPPED |  | Philips,Siemens |  | (0018, 0081) | Echo time | range | [45, 65] | 65
DWI,DWI_FLIPPED |  | GE |  | (0018, 0081) | Echo time | range | [45, 67.5] | 65
DWI,DWI_FLIPPED,B1MAP |  | Philips |  | (2001, 1018) | Number of slices | == | 13
DWI,DWI_FLIPPED |  | GE |  | (0021, 104F) | Number of slices | == | 13
B0 | B0MAP_PCASL | Philips |  | (0018, 0091) | Number of echoes | == | 2
B0 | B0MAP_PCASL | Philips,GE |  | (0018, 0081) | Echo time | contain | [4.0, 6.3]
B0 | B0MAP_PCASL | Siemens |  | (0018, 0081) | Echo time | contain | [4.0, 6.46]
B0 | B0MAP_PCASL |  |  | (0018, 1314) | Flip angle | or | [20,22]
B0 | B0MAP_PCASL |  |  | (0018, 0080) | TR | or | [130, 170] |  |  | 130 is correct but don't want to fail existing data
B0 | B0MAP_PCASL |  |  | (0018, 0095) | Bandwidth | range | [600, 900] | 787.5
MOLLI |  |  |  | (0018, 0081) | Echo time | range | [1, 2] | 1.68
MOLLI |  |  |  | (0018, 1314) | Flip angle | == | 20
MOLLI |  |  |  | (0018, 0080) | TR | range | [2, 4] | 3.4
MOLLI |  |  |  | (0018, 0095) | Bandwidth | range | [600, 900] | 787.5
MOLLI |  | GE |  | (0018, 1060) | Trigger time (first delay)
MOLLI |  |  |  | (0018, 1088) | Cardiac Frequency | == | 60
MOLLI |  | Philips |  | (0020, 9251) | RR Interval Time Nominal | == | 1000
MOLLI |  |  |  | (2005, 1572) | Inversion Delay Time | contain | [118, 201]
MOLLI |  |  |  | (2005, 1571) | Number of inversion Delays | == | 8
PC_LEFT,PC_RIGHT |  |  |  | (0018, 0081) | Echo time | range | [2, 4] | 3.3
PC_LEFT,PC_RIGHT |  |  |  | (0018, 1314) | Flip angle | == | 20
PC_LEFT,PC_RIGHT |  | GE,Philips |  | (0018, 0080) | TR | range | [4, 6] | 5.5
PC_LEFT,PC_RIGHT |  | Philips,Siemens |  | (0018, 0095) | Bandwidth | range | [350, 600] | 481.8
PC_LEFT,PC_RIGHT |  | GE |  | (0018, 0095) | Bandwidth | range | [320, 600] | 481.8
PC_LEFT,PC_RIGHT |  |  |  | (2005, 140f)/(0018,9069) | Parallel acquisition | == | 3
PC_LEFT,PC_RIGHT |  | Philips,Siemens |  | (0028, 0030) | Pixel spacing | == | [1.5, 1.5]
PC_LEFT,PC_RIGHT |  | GE |  | (0028, 0030) | Pixel spacing | == | [1.1, 1.1]
PC_LEFT,PC_RIGHT |  | Philips,GE,Siemens |  | (2001,1017) | Number of phases | == | 40
MTR |  | Philips |  | (2001,1018) | Number of slices | == | 1
MTR |  |  |  | (0018, 0081) | Echo time | == | 4.6
MTR |  |  |  | (0018, 1314) | Flip angle | == | 17 |  | WARNING
MTR |  |  |  | (0018, 0080) | TR | range | [50, 80] | 63 | WARNING
MTR |  |  |  | (0018, 0095) | Bandwidth | range | [600, 900] | 739.8 | WARNING
MTR |  | Philips |  | (2005, 140f)/(0018, 9069) | Parallel acquisition | == | 2
MTR |  | GE |  | (0043, 1083) | Parallel acquisition factors | == | [0.5, 1]
MTR |  | GE |  | (0019, 100a) | Number of slices | == | 1
MTR |  | Philips |  | (0018, 9020) | MagnetizationTransfer | == | OFF_RESONANCE
MDIXON |  | Philips |  | (2001,1018) | Number of slices | == | 31
MDIXON |  | Philips |  | (0018, 0091) | Number of echoes | or | [2, 6]
MDIXON |  |  |  | (0018, 1314) | Flip angle | range | [3, 9]
MDIXON |  |  |  | (0018, 0080) | TR | range | [4, 8]
MDIXON |  |  |  | (0018, 0095) | Bandwidth | range | [950, 2100] | 2034
T1MAP,T1MAP_FLIPPED, BASE,FAIR,PCASL | B0MAP_PCASL | Philips |  | (2001,1018) | Number of slices | == | 5
T1MAP,T1MAP_FLIPPED, BASE,FAIR,PCASL | B0MAP_PCASL | GE |  | (0021, 104F) | Number of slices | == | 5
T1MAP,T1MAP_FLIPPED, BASE,FAIR,PCASL | B0MAP_PCASL |  |  | (0018, 1314) | Flip angle | == | 90
T1MAP,T1MAP_FLIPPED, BASE,FAIR,PCASL | B0MAP_PCASL | Philips,Siemens |  | (0018, 0080) | TR | range | [3000, 6000] | 3000
T1MAP,T1MAP_FLIPPED |  | GE |  | (0018, 0080) | TR | range | [3000, 6000] | 6000
T1MAP,T1MAP_FLIPPED, BASE,FAIR,PCASL | B0MAP_PCASL |  |  | (0018, 0095) | Bandwidth | range | [1800, 2200] | 2049
T1MAP,T1MAP_FLIPPED, BASE,FAIR,PCASL | B0MAP_PCASL |  |  | (0018, 0081) | Echo time | range | [20, 35] | 29
T1MAP_FLIPPED |  |  |  | (0020, 0105) | Number of averages | == | 1
T1MAP | FLIPPED |  |  | (0020, 0105) | Number of averages | == | 20
BASE |  | Philips |  | (0020, 0105) | Number of averages | == | 2
PCASL_750,PCASL_500,PCASL_1000,FAIR_400,FAIR_750,FAIR_1000 |  | Philips |  | (0020, 0105) | Time points | == | 8
FAIR_1500,PCASL_1500 |  | Philips |  | (0020, 0105) | Time points | == | 25
AO |  |  |  | (0018, 0081) | Echo time | range | [2, 4] | 2.5
AO |  |  |  | (0018, 1314) | Flip angle | == | 15
AO |  | GE,Philips |  | (0018, 0080) | TR | range | [2, 4] | 3.8
AO |  |  |  | (0018, 0095) | Bandwidth | range | [750, 2000] | 1436 | WARNING
AO |  |  |  | (2005, 140f)/(0018, 9081) | Partial Fourier | == | NO
AO |  |  |  | (2005, 140f)/(0018, 9069) | Parallel acquisition | == | 2